## Contents
PCG60074373r13
This PCG is setup to retrieve the prices from the 'Local Price List' sheet
Copy prices from the Foss supplied 'Local Price List' sheet price list into this sheet.
OR enter local prices into the 'Local Price List' sheet
MilkoScan FT2'
MilkoScan FT2 UPGRADE'!A1
MSc FT2 Calibration Upgrades'!A1
Reagents, Accessories, Spares'!A1
Local Price List'

## MilkoScan FT2
To Contents
| Date issued: | 2020-10-16
PCG60074373r13
Tony Maddison
MilkoScan FT2
0
Total cost of configuration:
| 4404425 | MilkoScan FT2 Analyser
Start the order using this SAP BOM number and then make the required selections below. Includes Foss Integrator software and basiccalibrations: Milk calibrations: F, P, L, TS, SNF, FPD, TA, DEN, CA, FFA, U, CAS Cream calibrations: F, P, L, TS, SNF Concentrated milk calibrations: F, TS, SNF Infant formula calibrations: F, TS, SNF Liquid whey calibrations: F, P, L, TS
| 4404445 | MilkoScan FT2 Analyser UPGRADE Select S-Choice options below
Calibration options
Abbreviations for parameters: F=Fat, LF=Low Fat, P=Protein, L=Lactose, LL=Low Lactose, GLU=Glucose, SUC=Sucrose, FRU=Fructose, GAL=Galactose, TSU=Total Sugars, TS=Total Solids, SNF= Solids Non Fat, SALT=Salt, CA=CitricAcid, FP=Freezing Point, LA=Lactid Acid, U=Urea, CAS=Casein, MA=Malic Acid, AC=Acidity, LAC=Log(Acidity), DEN=Density, FFA=Fatty Acids, TA=Total Acidity.
You can order a complete calibration package including all the available calibrations or you can order the application packages individually.
| S158101 | Yoghurt & Fermented products applications - package includes: Yoghurt ( F, P, TS, SNF, L, GLU, FRU, SUC, TSU, LA ) Juice and Honey (FRU, GLU, SUC, TSU, MA, CA, TS ) Unfermented yoghurt (F, P, TS, SNF, L, GLU, FRU, SUC, TS) Quark (F, P, TS )
| S158102 | Dessert & Ice cream applications - package includes: Dessert ( F, P, TS, L, GLU, FRU, SUC, TSU ) Juice and honey (FRU, GLU, SUC, TSU, MA, CA, TS ) Sugar solutions (FRU, GLU, SUC, GAL, L, TSU)
| S158103 | Cheese & Whey applications - package includes: Cheese (F, P, TS, SALT ) Concentrated whey (P, TS, TA)
| S158104 | Total application package includes the above 3 pacakges: - Yoghurt & Fermented application package - Dessert & Ice cream application package - Cheese & Whey application package
| S158190 | Targeted Indentification: Raw Cow Milk Models Identifies adulterants in Raw Milk with the following targeted models: - Hydroxyproline - Sodium nitrite - Melamine - Maltodextrine - Cyanuric Acid - Formaldehyde - Sucrose - Bicarbonate
Software Options
| S158181 | Simple User Interface A simple operator interface to make daily routine analysis
| S158183 | LIMS option Possibility to set up connections for exporting result data to a laboratory software programme or to a file. Different formats are supported, XML, to text file format (csv), to a database (SQL)
| S158184 | Control Charts (Zero and S/I) Option to view and monitor graphs for zero-settings and S/I adjustments
| S158185 | Custom Job Type option Sends results acquired in the result list to a selected Excel sheet. The Excel sheet needs to follow a defined template. Calculations can be made in the Excel sheet and then automaticallly sent back to the result list to be shown
| S158186 | WinISI component transfer Converts Foss Integrator sample sets to WinISI sets and converts WinISI calbration files to Foss Integrator components and Mosaic prediction models
| S158187 | Re-prediction option Re-predicts samples with another measure profile. This is a pure re-calculation so the sample does not have to be re-measured.
| S158188 | Dilution Balance option For the dilution of samples with a balance is recommended for easy preparation of solid and semi-solid products.. The dilution factor obtained from the connected balance will automatically correct the predicted results. Interfacing with Mettler or Sartorius balances
| S158189 | Automatic Job Control option Possibility to let the sample ID control the analyser - it can be defined how the analyser should handle samples by looking at the sample ID
FTIR Calibration Development Option
This is a seperate program that can be installed on any computer for convenient use and so will not distrupt the operation of the FT2. Milkoscan FT2 instruments require Integrator version 1.6.0 or higher.
| S231501 | FTIR Calibrator PLS Development Develop calibrations for your own products or new measurment parameters using features such as MLR, PLS and GH components. No additional license is required to import PLS calibration into Foss Integrator as this is always included with the instrument.
| S231502 | Untargeted Abnormal Sample Screening Package Consists of: 1) FTIR Calibrator ASM Software Tools to develop and monitor Untargeted Screening Models. (Installation CD and License required to be requested after the software is installed from License@foss.dk ) 2) Foss Global Untargeted Raw Cows Milk Model. (Model is in the Component Installer on a CD or can be downloaded) 3) Integrator Import .asm models license. (Requested from License@foss.dk)
Product Configuration Guide

## MilkoScan FT2 UPGRADE
To Contents
| Date issued: | 2020-10-16
PCG60074373r13
Tony Maddison
MilkoScan FT2 UPGRADES
0
Total cost of configuration:
| 4404445 | MilkoScan FT2 Analyser UPGRADE Select S-Choice options below
Software Options
| S158181 | Simple User Interface A simple operator interface to make daily routine analysis
| S158183 | LIMS option Possibility to set up connections for exporting result data to a laboratory software programme or to a file. Different formats are supported, XML, to text file format (csv), to a database (SQL)
| S158184 | Control Charts (Zero and S/I) Option to view and monitor graphs for zero-settings and S/I adjustments
| S158185 | Custom Job Type option Sends results acquired in the result list to a selected Excel sheet. The Excel sheet needs to follow a defined template. Calculations can be made in the Excel sheet and then automaticallly sent back to the result list to be shown
| S158186 | WinISI component transfer Converts Foss Integrator sample sets to WinISI sets and converts WinISI calbration files to Foss Integrator components and Mosaic prediction models
| S158187 | Re-prediction option Re-predicts samples with another measure profile. This is a pure re-calculation so the sample does not have to be re-measured.
| S158188 | Dilution Balance option For the dilution of samples with a balance is recommended for easy preparation of solid and semi-solid products.. The dilution factor obtained from the connected balance will automatically correct the predicted results. Interfacing with Mettler or Sartorius balances
| S158189 | Automatic Job Control option Possibility to let the sample ID control the analyser - it can be defined how the analyser should handle samples by looking at the sample ID
FTIR Calibration Development Option
This is a seperate program that can be installed on any computer for convenient use and so will not distrupt the operation of the FT2. Milkoscan FT2 instruments require Integrator version 1.6.0 or higher.
| S231501 | FTIR Calibrator PLS Development Develop calibrations for your own products or new measurment parameters using features such as MLR, PLS and GH components. No additional license is required to import PLS calibration into Foss Integrator as this is always included with the instrument.
| S231502 | Untargeted Abnormal Sample Screening Package Consists of: 1) FTIR Calibrator ASM Software Tools to develop and monitor Untargeted Screening Models. (Installation CD and License required to be requested after the software is installed from License@foss.dk ) 2) Foss Global Untargeted Raw Cows Milk Model. (Model is in the Component Installer on a CD or can be downloaded) 3) Integrator Import .asm models license. (Requested from License@foss.dk)
Product Configuration Guide

## MSc FT2 Calibration Upgrades
To Contents
| Date issued: | 2020-10-16
PCG60074373r13
Tony Maddison
MilkoScan FT2 Calibration Upgrades
0
Total cost of configuration:
| 4404555 | MilkoScan FT2 Analyser
Calibration options
Abbreviations for parameters: F=Fat, LF=Low Fat, P=Protein, L=Lactose, LL=Low Lactose, GLU=Glucose, SUC=Sucrose, FRU=Fructose, GAL=Galactose, TSU=Total Sugars, TS=Total Solids, SNF= Solids Non Fat, SALT=Salt, CA=CitricAcid, FP=Freezing Point, LA=Lactid Acid, U=Urea, CAS=Casein, MA=Malic Acid, AC=Acidity, LAC=Log(Acidity), DEN=Density, FFA=Fatty Acids, TA=Total Acidity.
You can order a complete calibration package including all the available calibrations or you can order the application packages individually.
| S158101 | Yoghurt & Fermented products applications - package includes: Yoghurt ( F, P, TS, SNF, L, GLU, FRU, SUC, TSU, LA ) Juice and Honey (FRU, GLU, SUC, TSU, MA, CA, TS ) Unfermented yoghurt (F, P, TS, SNF, L, GLU, FRU, SUC, TS) Quark (F, P, TS )
| S158102 | Dessert & Ice cream applications - package includes: Dessert ( F, P, TS, L, GLU, FRU, SUC, TSU ) Juice and honey (FRU, GLU, SUC, TSU, MA, CA, TS ) Sugar solutions (FRU, GLU, SUC, GAL, L, TSU)
| S158103 | Cheese & Whey applications - package includes: Cheese (F, P, TS, SALT ) Concentrated whey (P, TS, TA)
| S158104 | Total application package includes the above 3 pacakges: - Yoghurt & Fermented application package - Dessert & Ice cream application package - Cheese & Whey application package
| S158190 | Targeted Indentification: Raw Cow Milk Models Identifies adulterants in Raw Milk with the following targeted models: - Hydroxyproline - Sodium nitrite - Melamine - Maltodextrine - Cyanuric Acid - Formaldehyde - Sucrose - Bicarbonate
Product Configuration Guide

## Reagents, Accessories, Spares
To Contents
| Date issued: | 2020-10-16
PCG60074373r13
Tony Maddison
MilkoScan FT2
Reagents, Accessories, Spares
Reagents
| 60083360 | FTIR Equalizer Reagent 6 bottles, 6 x 100 mls PLASTIC BOTTLES For use during 3-6 months. Standardisation is required at instrument installation and then after every 2 to 4 weeks depending on the number and type of samples analysed. Transport temperature not critical. Storage for use at 2-8 C
| 60093992 | MSc Zero Concentrate Replaces MSc Zero, 60077631 Used for rinsing and Zero the FT2 flow system 50 sachets; each for 5 liters. For 3 to 6 months use Shelf life minimum 1 year from shipping Use within 7 days when mixed
| 60083967 | MSc Wash: This Replaces S-470 50 Sachets; each for 5 liters. NEW improved cleaning solution and with a purple color for easy identification Used for daily cleaning by the instruments automatic cleaning cycle.
| 60093996 | MilkoScan Flush Concentrate Used as an alternative to the MSc Wash in cases where hogh cuvette wear is reported. Addiotnal daily cleaning is still required with MSc Wash. Box containing 50 bottles, each 50 ml Detained transport conditions: Temperature not critical
| 60070869 | MSc-DeScaler 10 Sachets; each for 1 liters. Used for daily/weekly 'Extra-Clean' procedure. Keeps the flow system clean from mineral build up. Shelf life minimum 1 year from shipping Use within 7 days when mixed
| 541011 | FossClean Enzymatic cleaning detergent Package of 2 bottles. For removing protein build-up in flow system. Approx. duration: 2 months. Detained transport conditions: Keep cool (2-8ºC)
| 67165 | Antifoam Y-30 Air bubble remover for products with a foaming texture. Detained transport conditions: Temperature not critical
Consumables
| 399436 | Sample Cups, 30 ml, with caps 750 pcs/pkg
Spare Parts, Consumables & Accessories
| 482612 | Yoghurt Filtration Unit: Can Also be purchased on Part Number: 482612 (S158167) A Stainless Steel filter, 300µm, to remove fruit and larger particles from the yoghurt and Ice Cream samples.
| 60015936 | FOSS PM.Kit, MSCFT2
| 60035281 | FOSS Cus.Tool.Kit, MSCFT2
| 60035689 | FOSS Cus.Mtn.Kit, MSCFT2
| 60088703 | FT 2, Cuvette (SP)
| 60045456 | In-Line Filter, Mesh, 50µm, 10 pieces For products such as milk and cream
| 60053543 | In-Line Filter, Screen, 24µm, 10 pieces For products with added solids such as cocoa milk and flavoured milks
| 60049648 | Inline Filter Assembly Complete .FT2
| 60077183 | Milk Filter Bag 50 micron Nylon Mesh Filter bag, reusable and washable For filtering liquid samples containing particles such as raw whey
| 60079482 | Milk Filter Bag 250 micronNylon Mesh Filter bag, reusable and washable For filtering liquid and thick samples containing particles such as yoghurt or ice cream mix
| 60082536 | Pipette Filter Holder. Optional sample pipette filter similar to the MilkoScan Minor Uses replaceable 150µm filters 1015815
| 1015815 | Pipette Filters (12 pieces) 150 µm filters for the above filter holder
| 1015263 | Container Clean 5L with labels
| 1015262 | Container Zero 5L with labels
Product Configuration Guide